## Slide 1: 3sided Pyramid @ PAPYRUS Run 5th-9th August Data from 8th August
Muskan Shinde

## Slide 2
First image
after increasing the pupil sampling to 35 pixels

## Slide 3
Finding a good flat
After repeated closing the loop on first 5-10 modes and calibrations

## Slide 4
Push-Pull Calibrations
Calibration procedure:
Took 100 measurements of push and pull for each of the first 150 modes

## Slide 5
Valid Pixels Mask

## Slide 6
Problem in the calibration!
The full images during calibrations were NOT saved, only the masked images were saved.
And the mask was changed after.
Zeros added in the interaction matrix

## Slide 7
There are some zeros in the the interaction matrix

## Slide 8
Applying Turbulence through DAO GUI
Applied seeing | Estimated seeing
4.0″ | 2.40″
3.0″ | 1.92″
2.0″ | 1.40″
1.0″ | 0.80″
Plate scale = 40 mas/pix

## Slide 9
Turbulence speed: 500 Hz
3s pyramid speed: 500 Hz
4s pyramid speed: 500Hz
Closing loop
on injected turbulence
@ 500Hz

## Slide 10
RMS residual modes measured by the pyramid
Why are the rms residual of the 3-sided pyramid higher than 4-sided pyramid?
4-sided pyramid is underperforming and what see are not true residuals
Calibration used was with modulation
Loop gain not set according to compensate optical gains

## Slide 11
Closing loop
on injected turbulence
@ 1000Hz
Turbulence speed: 500 Hz
3s pyramid speed: 1000 Hz
4s pyramid speed: 500Hz
- What is 500 Hz turbulence?
- Is 3s Pyramid at 1000Hz observing the same frame twice?

## Slide 12
Seeing | OL | CL (3s, 500 Hz) | CL (3s, 1000 Hz) | CL (4s, 500 Hz)
1″ | 0.0218 | 0.2554 | 0.2954 | 0.1147
2″ | 0.0070 | 0.1806 | 0.2295 | 0.0504
3″ | 0.0040 | 0.1200 | 0.1337 | 0.0239
4″ | 0.0026 | 0.1416 | 0.1690 | 0.0189
Encircled Enery
is 3x3 pixels area around the center

## Slide 13
Missed frames
Seeing | CL (3s, 1000 Hz)/CRED3 |  | CL (3s, 500 Hz)/CRED3 |  | CL (4s, 500 Hz)/OCAM | 
 | Missed frames | Duplicate frames | Missed frames | Duplicate frames | Missed frames | Duplicate frames
1″ | 1339.2 | 71 | 22 | 11 | 187.5 | 32
2″ | 191.3 | 158 | 59.5 | 56 | 193 | 18
3″ | 72.9 | 13 | 38.5 | 34 | 8 | 3
4″ | 320.2 | 287 | 106.5 | 98 | 49.5 | 17
avg | 480.9 | 132.3 | 56.6 | 49.8 | 109.5 | 17.5
missed frames & duplicate frame during the day
~ 610 out 5000 frames
~ 12.2%
~106 out 5000 frames
~ 2.1%
~127 out 5000 frames
~ 2.5%

## Slide 14
On Sky tests
3s pyramid speed: 1000 Hz
4s pyramid speed: 500Hz
Why the 3s pyramid could not work on sky?
NIR pupil conjugation issue?
NIR pupil vignetting?
Too many missed frames?

## Slide 15
Missed frames
missed frames & duplicate frame during the night
CL (3s, 1000 Hz)/CRED3 |  | CL (4s, 500 Hz)/OCAM | 
Missed frames | Duplicate frames | Missed frames | Duplicate frames
438.8 | 314 | 150 | 13
~ 753 out 5000 frames
~ 15%
~ 163 out 5000 frames
~ 3.3%
Most of the times there was a missed frame after a duplicate frame

## Slide 16
To Do for next time:
Fix the saving of full calibration with all the frames
Complete telemetry with slopes and frame counter
Record data while closing the loop
Reduce the missed frames & jitter - put code in DAO
Synchronize taking data from 3s and 4s pyramid
Synchronize with turbulence
PSF centering algorithm
Make a workbook to take the data
Optical gains for 4s pyramid
Analysing the data from 7th August
Internal turbulence injection studying